# Operational Procedures – Intoxilyzer 8000

1.	All results are reported to three decimal places in g/210L.  The instrument must display READY MODE prior to beginning the breath test.  Push the START TEST button to begin the breath test:

- Enter Breath Test Operator last name, first name and middle initial at USER prompts.

- DATE/TIME.  Verify the displayed date and time.  Change if necessary.  Press ENTER.

- LAST AGENCY INSPECTION DATE.  Verify the agency inspection date displayed.  Press ENTER.

- CYLINDER LOT#.  Verify the dry gas standard cylinder lot number.  Change if necessary.  Press ENTER.

- EXPIRATION DATE.  Verify the dry gas standard cylinder expiration date.  Change if necessary.  Press ENTER.

- OBSERVATION PERIOD BEGAN.  Enter the time the observation period began (at least 20 minutes).

- SWIPE DL OR PRESS ENTER.  Either swipe the subject’s driver license or identification card or press ENTER.  Enter applicable information as prompted.

- DIAGNOSTICS CHECK.  The result must be OK.

- AIR BLANK.  The result must be 0.000.

- CONTROL TEST.  The result must be between 0.075 and 0.085, inclusive.

- AIR BLANK.  The result must be 0.000.

- PROVIDE SAMPLE NOW.  Have the subject provide a breath sample into the instrument.

- AIR BLANK.  The result must be 0.000.

- PLEASE WAIT.  The instrument will countdown the time remaining for the wait period.

- AIR BLANK.  The result must be 0.000.

- PROVIDE SAMPLE NOW.  Have the subject provide a breath sample into the instrument.

- AIR BLANK.   The result must be 0.000.

Note:	If there is no 0.020 g/210L agreement between first and second breath samples, the instrument will automatically request a third breath sample as follows:

- PLEASE WAIT.  The instrument will countdown the time remaining for the wait period.

- AIR BLANK.  The result must be 0.000.

- PROVIDE SAMPLE NOW.  Have the subject provide a breath sample into the instrument.

- AIR BLANK.  The result must be 0.000.

- CONTROL TEST.  The result must be between 0.075 and 0.085, inclusive.

- AIR BLANK.  The result must be 0.000.

- DIAGNOSTICS CHECK.  The result must be OK.

2.	If an external printer is used, FDLE/ATP Form 38 – Breath Alcohol Test Affidavit - Intoxilyzer 8000 will be automatically printed containing all the results.  If no external printer is used, a printout slip containing all the results will be automatically printed.

- Complete FDLE/ATP Form 38 – Breath Alcohol Test Affidavit - Intoxilyzer 8000.